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	Get exactly what you need with the totally reengineered X-50. It’s the highest performing XRF on the market that features a cost effective SDD detector. A little beefier than SciAps sleeker models, this cost-effective XRF still offers best-in-class analytical performance and speed for its detector platform, operating at rates 2X or higher than other brands. SciAps powerful, miniaturized X-ray tube combined with highly advanced internal geometry yields fast, precise results. Analyze all of your precious metals, quickly and easily the SciAps X-50.
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	The classic model for many applications at a great value.
	

	
	Premium X-ray hardware for reliable handling
	

	
	Optimal performance on precious metals like Au, Ag, Pt and Pd
	

	
	Fast, precise results
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	X-50
	

	
	XRF
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	Unmatched Quality at a Value Price
	

	
	The X-50 is great for fast, precise results on Au, Ag, Pt, and Pd concentrations, making valuable metals even more profitable. We’ve reengineered the classic detector technology with revamped electronics, improved ergonomics, updated calibration and the latest software and user interface. And we’ve equipped it with more features, such as a built-in, high-resolution camera for sample viewing and a macro-camera for photo-documentation or 2D/3D barcode reading and storage.
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	Standard appliction package
	

	
	The X-50 operates using the same advanced X-ray tube technology as other SciAps X-Series models (operating at 40 kV max) for testing, that includes Ti, V, Cr, Mn, Fe, Co, Ni, Cu, Zn, Ir, Pt, Au, Pb, Bi, Zr, Mo, Ru, Rh, Pd, Ag, Cd, Sn and Sb. More elements can be added upon request.
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	Android and data management
	

	
	Operates on Android OS with the feel of a smartphone. Using Bluetooth, Wi-Fi and USB, users can print, email, and connect to virtually any information system for real-time data. Onboard macro camera allows for photo-documentation, and Bluetooth label printer provides instant hard copy labels.
	

	
	Use SciAps Test Station to analyze small pieces in benchtop mode. Features an interlocking lid for your protection and super stable base to keep samples positioned correctly.
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	For more information, or to schedule a demonstration:
	

	
	If applicable, please add: 
LOCAL COMPANY NAME
LOCAL WEBSITE
LOCAL PHONE
	

	(BACK)
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	A classic model for many applications at a great value.
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	Weight
	

	
	3.1 lbs. ( 1.40kg) with battery
	

	
	Dimensions
	

	
	9.38in (238mm) x 11.15in (283mm) x 3.34in (84mm)
	

	
	Excitation Source
	

	
	4 W, 40kv Rh on alloys 50kv Au on other apps.
	

	
	Detector
	

	
	7 mm² SDD, silicon drift detector, (active area), 170 eV resolution FWHM at 5.95 Mn K-alpha line.
	

	
	Available Apps
	

	
	Alloy analysis, Precious Metals, Soil, Mining, Car Cats. New apps are added regularly, please check with company or website.
	

	
	X-ray Filtering
	

	
	Single primary beam filter for alloy, mining, precious metals. Multiple filters for soil, RoHS, car catalysts and some other apps.
	

	
	Environmental Temperature Range
	

	
	10F to 130F at 25% (-12.2C to 54.44C) duty cycle
	

	
	Analytical Range
	

	

	23 elements standard, specific elements vary by app. Ti, V, Cr, Mn, Fe, Co, Ni, Cu, Zn, Ir, Pt, Au, Pb, Bi, Zr, Mo, Ru, Rh, Pd, Ag, Cd, Sn and Sb. Additional elements may be added upon user request. Precious Metals app is 23 elements standard.
	

	
	Processing Electronics and Host Processing
	

	
	1.2 GHz quad ARM Cortex A53 64/32-bit; RAM: 2 GB LP-DDR3; Storage: 16 GB eMMC (storage).
	

	
	Pulse Processor
	

	
	12 bit with digitization rate of 80 MSPS 8K channel MCA USB 2.0 for high-speed data transfer to host processor. Digital filtering implemented in FPGA for high throughput pulse processing, 20 nS - 24 uS peaking time.
	

	
	Power
	

	
	On-board rechargeable Li-ion battery, rechargeable inside device or with external charger, AC power, hot-swap capability (60 s max swap time).
	

	
	Display
	

	
	3.5-inch(88.9mm) color capacitive touchscreen — 400 MHz Qualcomm Adreno 306 2D/3D graphics accelerator
	

	
	Sample Viewing
	

	
	Internal camera for viewing sample before and during analysis for proper sample alignment. Second macro-camera for scanning QR or barcodes and for photo-documentation and report generation.
	

	
	Comms/Data Transfer
	

	
	Wi-Fi, Bluetooth, USB connectivity to most devices, including SciAps Profile Builder PC software.
	

	
	Calibration
	

	
	Fundamental parameters.
	

	
	Calibration Check
	

	
	External 316 stainless check standard for calibration verification and energy scale validation.
	

	
	Security
	

	
	Password protected usage (user level) and internal settings (admin).
	

	
	Dual Cameras
	

	
	Internal high-resolution camera for sample viewing, welds, etc. Macrocamera for photo documentation, reading and storing 2D/3D barcodes and QR codes.
	

	
	Regulatory
	

	
	CE, RoHS, USFDA registered, Canada RED Act.
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	If applicable, please add: 
LOCAL COMPANY NAME
LOCAL WEBSITE
LOCAL PHONE
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	YouTube.com/SciAps
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Unmatched Quality at a Value Price

The X-50 is great for fast, precise results on Au, Ag, Pt, and Pd concentrations,
making valuable metals even more profitable. We’ve reengineered the classic
detector technology with revamped electronics, improved ergonomics, updated
calibration and the latest software and user interface. And we’ve equipped it with
more features, such as a built-in, high-resolution camera for sample viewing and a
macro-camera for photo-documentation or 2D/3D barcode reading and storage.
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Standard appliction package

The X-50 operates using the same advanced X-ray tube
technology as other SciAps X-Series models (operating at 40
kV max) for testing, that includes Ti, V, Cr, Mn, Fe, Co, Ni, Cu,
Zn, Ir, Pt, Au, Pb, Bi, Zr, Mo, Ru, Rh, Pd, Ag, Cd, Sn and Sb. More
elements can be added upon request.
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Android and data management

Operates on Android OS with the feel of a smartphone. Using
Bluetooth, Wi-Fi and USB, users can print, email, and connect
to virtually any information system for real-time data. On-
board macro camera allows for photo-documentation, and
Bluetooth label printer provides instant hard copy labels.

Use SciAps Test Station to analyze small pieces in benchtop mode.
Features an interlocking lid for your protection and super stable base
to keep samples positioned correctly.
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schedule a demonstration:
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3.1 Ibs. (1.40kg) with battery

9.38in (238mm) x 11.15in (283mm) x 3.34in (84mm)

4 W, 40kv Rh on alloys 50kv Au on other apps.

7 mm? SDD, silicon dfrift detector, (active area), 170 eV resolution FWHM
at 5.95 Mn K-alpha line.

Alloy analysis, Precious Metals, Soil, Mining, Car Cats. New apps are
added regularly, please check with company or website.

Single primary beam filter for alloy, mining, precious metals. Multiple
filters for soil, RoHS, car catalysts and some other apps.

10F to 130F at 25% (-12.2C to 54.44C) duty cycle

23 elements standard, specific elements vary by app. Ti, V, Cr, Mn, Fe, Co,
Ni, Cu, Zn, Ir, Pt, Au, Pb, Bi, Zr, Mo, Ru, Rh, Pd, Ag, Cd, Sn and Sb.
Additional elements may be added upon user request. Precious Metals
app is 23 elements standard.

12 GHz quad ARM Cortex A53 64/32-bit; RAM: 2 GB LP-DDR3; Storage:
16 GB eMMC (storage).

12 bit with digitization rate of 80 MSPS 8K channel MCA USB 2.0 for
high-speed data transfer to host processor. Digital filtering implemented
in FPGA for high throughput pulse processing, 20 nS - 24 uS peaking
time.

On-board rechargeable Li-ion battery, rechargeable inside device or with
external charger, AC power, hot-swap capability (60 s max swap time).

3.5-inch(88.9mm) color capacitive touchscreen — 400 MHz Qualcomm
Adreno 306 2D/3D graphics accelerator

Internal camera for viewing sample before and during analysis for proper
sample alignment. Second macro-camera for scanning QR or barcodes

and for photo-documentation and report generation.

Wi-Fi, Bluetooth, USB connectivity to most devices, including SciAps
Profile Builder PC software.

Fundamental parameters.

External 316 stainless check standard for calibration verification and
energy scale validation.

Password protected usage (user level) and internal settings (admin).
Internal high-resolution camera for sample viewing, welds, etc.
Macrocamera for photo documentation, reading and storing 2D/3D

barcodes and QR codes.

CE, RoHS, USFDA registered, Canada RED Act.
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Get exactly what you need with the totally reengineered X-50. It's the
highest performing XRF on the market that features a cost effective SDD
detector. A little beefier than SciAps sleeker models, this cost-effective XRF
still offers best-in-class analytical performance and speed for its detector
platform, operating at rates 2X or higher than other brands. SciAps
powerful, miniaturized X-ray tube combined with highly advanced internal
geometry yields fast, precise results. Analyze all of your precious metals,
quickly and easily the SciAps X-50.
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